# Protocol: Measuring the Attenuation Coefficient of Hydrogels

As of October 27, 2026

## 1. Purpose and principle

This protocol measures the ultrasonic attenuation coefficient α of a hydrogel, in dB/cm and dB/cm/MHz, by pulse-echo (Method A) or substitution (Method B). An optional optical method is in section 7.

A wave's amplitude decays exponentially as it travels through the gel:

Ax=A0 e−αx

So for a path length d through the gel:

αdB/cm=20d log10​A0A=8.686 αNp/cm

Attenuation combines absorption and scattering. In hydrogels it depends mainly on polymer concentration, crosslink density, porosity, temperature and frequency.

## 2. Materials and equipment

| Item | Specification / purpose |
| --- | --- |
| Ultrasonic pulser-receiver | Broadband pulse; adjustable gain and damping |
| Transducer(s) | Immersion, 1–10 MHz centre frequency; two matched units for Method B |
| Digital oscilloscope or DAQ | ≥ 100 MS/s sampling (≥ 10× the transducer frequency); waveform export |
| Water tank | Degassed, deionised water; transducer holder and sample stage |
| Reflector plate | Polished steel or glass, flat (Method A) |
| Thermometer | ±0.1 °C; record water and sample temperature |
| Calipers or micrometer | ±0.01 mm, for sample thickness |
| Moulds | Flat-faced, at least 3 thicknesses (e.g. 3, 5, 8 mm), diameter > 2× transducer face |
| Vacuum desiccator or sonicator | Degassing the precursor solution |
| Balance | ±0.1 mg, for density (ρ) and swelling |
| Hydrogel precursors | Polymer, crosslinker, initiator; record lot numbers |
| Software | Python/MATLAB for FFT and fitting |

## 3. Sample preparation

Bubbles and uneven thickness are the main sources of error, so degas every precursor and cast flat, parallel faces.

- Prepare the precursor at the target concentration (% w/v). Record the polymer, crosslinker and initiator amounts.

- Degas the precursor for 10–15 min under vacuum or in a sonicator until no bubbles rise.

- Pour slowly down the mould wall to avoid trapping air. Cast at least three thicknesses per formulation, n ≥ 3 samples each.

- Cure by your standard method (UV, thermal, ionic). Record the time, light intensity or temperature.

- Equilibrate in degassed water or buffer for 24 h, or until mass is stable, at the measurement temperature.

- Measure thickness d at 5 points per sample, without compressing the gel, and record the mean ± SD. Reject samples whose faces differ by more than 2%.

- Measure density ρ (mass ÷ volume, or by buoyancy), which you need for the acoustic impedance.

- Keep samples hydrated and measure them within the same day.

## 4. Method A: pulse-echo

One transducer sends a pulse and receives the echoes from the gel's top surface and from the reflector beneath it.

- Fill the tank with degassed water and let it reach a stable temperature (e.g. 25.0 ± 0.2 °C).

- Mount the transducer facing down, normal to the reflector plate. Position it so the gel sits in the far field, beyond the transducer's near-field distance N = D²/(4λ).

- Reference: with no sample, record the echo from the reflector, A_ref. Tilt the transducer until this echo is at its maximum.

- Sample: place the gel flat on the reflector without trapping air. Do not change the gain, the damping or the transducer height.

- Record two echoes:

  - A₁, from the water–gel top surface

  - A₂, from the gel–reflector interface, after crossing the gel twice

- Average at least 64 waveforms per acquisition. Save the full time-domain signals, not just the peak values.

- Measure the time between the two echoes, Δt. The sound speed is c = 2d/Δt.

- Repeat at 3 positions per sample and on every thickness.

- Record the water temperature at the start and end of each run.

## 5. Method B: substitution (through-transmission)

Two coaxial transducers face each other in water. You compare the signal with and without the gel in the beam path.

- Align the transmitter and receiver coaxially, with the receiver in the transmitter's far field. Adjust the alignment until the received signal is at its maximum.

- Reference: record the signal through water only, A_ref(t).

- Sample: insert the gel perpendicular to the beam, centred, in a holder that leaves both faces in water. Keep all settings unchanged.

- Record A_s(t) and the arrival-time shift Δt relative to the reference.

- Compute the sound speed in the gel from the time shift:

cs=d d/cw−Δt

Here c_w is the speed of sound in water at the measured temperature, and Δt > 0 means the gel signal arrives earlier.

- Average at least 64 waveforms. Repeat at 3 positions and on every thickness.

- Re-record the water reference at the end to check for drift. Repeat the run if the two references differ by more than 2%.

## 6. Data analysis

Report α in dB/cm at each frequency, and fit α₀ (dB/cm/MHz) and n over the transducer's −6 dB bandwidth.

### 6.1 Acoustic impedance and interface coefficients

For each material, Z = ρ·c. With water (w), gel (s) and reflector (r):

Rws=Zs−ZwZs+Zw,  TwsTsw=1−Rws2,  Rsr=Zr−ZsZr+Zs

Hydrogels have Z close to water's, so R_ws is small and the top-surface echo A₁ is weak. That makes the reflector-referenced form below more reliable than using A₁/A₂.

### 6.2 Method A (pulse-echo)

Use the reflector echo without the sample (A_ref) and the back-wall echo with the sample (A₂):

αs=12d ln​ArefA2⋅1−Rws2 RsrRwr+αw Np/cm

The α_w term accounts for the water path the gel replaces; it is negligible below about 5 MHz. Multiply by 8.686 for dB/cm.

### 6.3 Method B (substitution)

αs=αw+1d ln​ArefAs⋅TwsTsw Np/cm

Water attenuation is about 0.0022 · f² dB/cm at 20 °C, with f in MHz.

### 6.4 Multi-thickness check (recommended)

Plot ln(A₂ or A_s) against path length (2d for Method A, d for Method B) across your three thicknesses. The slope is −α, and the interface losses fall into the intercept. This removes the need for exact R and T values, so use it as the primary result when the plot is linear (R² > 0.98).

### 6.5 Frequency-dependent attenuation

- Gate each echo with a Hann or Tukey window of the same length.

- Take the FFT and get the magnitude spectra |S_ref(f)| and |S_s(f)|.

- Apply the equations above at each frequency inside the −6 dB bandwidth.

- Fit α(f) = α₀ · fⁿ by linear regression of log α against log f.

- For focused or small transducers, apply a diffraction correction, or use the multi-thickness method, which largely cancels it.

import numpy as np

def spectrum(sig, fs, t0, t1):
    seg = sig[int(t0*fs):int(t1*fs)] * np.hanning(int(t1*fs) - int(t0*fs))
    f = np.fft.rfftfreq(len(seg), 1/fs)
    return f, np.abs(np.fft.rfft(seg))

# Method B, Np/cm -> dB/cm; d in cm, fs in Hz
f, S_ref = spectrum(ref, fs, t0_ref, t1_ref)
_, S_s = spectrum(samp, fs, t0_s, t1_s)
f_MHz = f / 1e6
alpha_w = 0.0022 * f_MHz**2                       # dB/cm
alpha = 8.686 / d * np.log((S_ref / S_s) * TwsTsw) + alpha_w

band = (f_MHz > f_lo) & (f_MHz < f_hi)            # -6 dB bandwidth
n, log_a0 = np.polyfit(np.log(f_MHz[band]), np.log(alpha[band]), 1)
alpha0 = np.exp(log_a0)                           # dB/cm/MHz^n

## 7. Optional: optical attenuation by UV-vis

Use this method when you need light penetration, for example to judge photocuring depth or transparency.

- Cast the gel directly in a cuvette, or cut slabs of known path length d (1–10 mm).

- Blank the spectrophotometer with a cuvette of water or buffer.

- Record the absorbance A from 300 to 800 nm, or at your curing wavelength (e.g. 365 or 405 nm).

- Calculate the attenuation coefficient:

μ=lnI0/Id=2.303 Ad cm−1

- The penetration depth is 1/μ.

In turbid gels the result combines absorption and scattering, so report it as attenuation, not absorption. Use at least two path lengths to confirm A rises linearly with d.

## 8. Quality control and reporting

A result is valid only if the water references are stable, the samples are bubble-free and the multi-thickness fit is linear.

- ☐ Validate the setup on a reference material of known α (e.g. a commercial tissue-mimicking phantom, or a 3% agar gel) before testing new gels.

- ☐ Water temperature stayed within ±0.2 °C during each run.

- ☐ Start and end water references agree within 2%.

- ☐ No visible bubbles or delamination; any sample with a bubble in the beam path is discarded.

- ☐ Signals neither clip nor sit in the noise floor (SNR > 20 dB).

- ☐ At least 3 samples × 3 positions per formulation; report mean ± SD.

- ☐ Multi-thickness fit R² > 0.98.

Report for each formulation: polymer and concentration (% w/v), crosslinker ratio, cure conditions, swelling state, temperature, transducer centre frequency and bandwidth, method (A or B), d, ρ, c, α at the centre frequency (dB/cm), α₀ (dB/cm/MHz) and n.

Data sheet (one row per sample position)

| Sample ID | Formulation | Thickness d (mm) | Density ρ (g/cm³) | Temp (°C) | Method | Sound speed c (m/s) | α at f_c (dB/cm) | α₀ (dB/cm/MHz) | n | Notes |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |